## DisplayEval
Display Evaluation
Procedure: Measure the luminance response of the display, starting with the minimum luminance (gray level 0) and ending with the maximum luminance (gray level 255 on 8-bit systems). Measure the MLD uniformity in the center and either the four corners or four edges. Measure the LUDM uniformity at the center, edges, and corners. The uniformity measurement can be done at any luminance, but is typically performed at gray level 30, 120, or 210. At each measured gray level, also record the chromaticity coordinates (u',v') for the TG196 analysis. | [5 empty] | Current |  |  |  | Previous |  |  |  | GSDF | [12 empty] | Color
[5 empty] | GL | L | u' | v' | Prev GL | L | u' | v' | Prev L' | L' | JND | JNDgsdf | Lgsdf | dJND/GL | relErr | (dL/L)/JND | dL/L dcm | relErr | dcm+10% | dcm-10% | dcm+20% | dcm-20% | u' | v' | ΔD65 | CCT (K) | Δ1 | Δ2
[5 empty] | 0 | 1.32 | 0.2061 | 0.4802 | 0 |  |  |  |  | 1.42 | 86.9103 | 86.9103 | 1.4195 | [9 empty] | 0.1997 | 0.49181 | 0.023558 | 5327.8328 | 0.0077488
[5 empty] | 15 | 2.6 | 0.20594 | 0.48003 | [5 empty] | 2.7 | 120.8807 | 122.4467 | 2.7708 | 2.2647 | -0.044068 | 0.017485 | 0.01815 | -0.036638 | 0.019965 | 0.016335 | 0.02178 | 0.01452 | 0.1997 | 0.49181 | 0.023558 | 5327.8328 | 0.0077488 | 0
[5 empty] | 30 | 4.55 | 0.20551 | 0.48343 | [5 empty] | 4.65 | 156.118 | 157.9832 | 4.7734 | 2.3492 | -0.0084178 | 0.014931 | 0.014939 | -0.00050737 | 0.016433 | 0.013445 | 0.017927 | 0.011951 | 0.1997 | 0.49181 | 0.023558 | 5327.8328 | 0.0077488 | 0
[5 empty] | 45 | 7.31 | 0.1997 | 0.49181 | [5 empty] | 7.41 | 191.494 | 193.5196 | 7.5981 | 2.3584 | -0.0045155 | 0.01288 | 0.01285 | 0.0023081 | 0.014135 | 0.011565 | 0.01542 | 0.01028 | 0.1997 | 0.49181 | 0.023558 | 5327.8328 | 0.0077488 | 0
Evaluation Information |  |  |  |  | 60 | 11.09 | 0.19928 | 0.49274 | [5 empty] | 11.19 | 226.9246 | 229.0561 | 11.4589 | 2.362 | -0.0029806 | 0.011438 | 0.011402 | 0.0031438 | 0.012542 | 0.010262 | 0.013682 | 0.0091214 | 0.19928 | 0.49274 | 0.024454 | 5315.4902 | 0.0067574 | 0.0010142
Display Make |  | Luminance Meter | Evaluation Std | Report 270 | 75 | 16.37 | 0.19873 | 0.49574 | [5 empty] | 16.47 | 263.6933 | 264.5926 | 16.6233 | 2.4512 | 0.034675 | 0.010743 | 0.01035 | 0.037976 | 0.011385 | 0.0093152 | 0.01242 | 0.0082802 | 0.19873 | 0.49574 | 0.027425 | 5240.903 | 0.0042552 | 0.0030502
Display Model |  | Calibration Date | Display Type |  | 90 | 23.02 | 0.19776 | 0.49693 | [5 empty] | 23.12 | 298.7243 | 300.129 | 23.4271 | 2.3354 | -0.014222 | 0.0094535 | 0.0095609 | -0.011231 | 0.010517 | 0.0086048 | 0.011473 | 0.0076487 | 0.19776 | 0.49693 | 0.028601 | 5246.3968 | 0.0028188 | 0.0015349
Display Serial |  | Illuminance Meter | Date |  | 105 | 31.78 | 0.19913 | 0.49459 | [5 empty] | 31.88 | 334.1999 | 335.6655 | 32.2905 | 2.365 | -0.0017126 | 0.0089639 | 0.0089529 | 0.0012238 | 0.0098482 | 0.0080577 | 0.010744 | 0.0071624 | 0.19913 | 0.49459 | 0.026296 | 5260.0142 | 0.005283 | 0.0027099
Display DOM |  | Calibration Date | Evaluator |  | 120 | 43.09 | 0.1981 | 0.49698 | [5 empty] | 43.19 | 369.6624 | 371.202 | 43.7402 | 2.3642 | -0.0020816 | 0.0084791 | 0.0084754 | 0.00043953 | 0.009323 | 0.0076279 | 0.010171 | 0.0067803 | 0.1981 | 0.49698 | 0.028653 | 5229.487 | 0.0031077 | 0.0026019
[5 empty] | 135 | 57.62 | 0.19674 | 0.49524 | [5 empty] | 57.72 | 405.168 | 406.7384 | 58.4356 | 2.367 | -0.00086929 | 0.0081038 | 0.0080945 | 0.0011481 | 0.0089039 | 0.007285 | 0.0097134 | 0.0064756 | 0.19674 | 0.49524 | 0.026932 | 5349.3779 | 0.0032752 | 0.0022075
Qualitative Components |  |  | Luminance Analysis |  | 150 | 76.37 | 0.19732 | 0.49576 | [5 empty] | 76.47 | 441.0114 | 442.2749 | 77.2024 | 2.3896 | 0.0086372 | 0.0078639 | 0.0077869 | 0.0098844 | 0.0085656 | 0.0070082 | 0.0093443 | 0.0062295 | 0.19732 | 0.49576 | 0.027437 | 5305.1527 | 0.0031721 | 0.00077674
Visible Inspection |  | Pass | Lamb (cd/m2) | 0.1 | 165 | 99.81 | 0.1977 | 0.49634 | [5 empty] | 99.91 | 476.2433 | 477.8113 | 101.0749 | 2.3488 | -0.008569 | 0.0074794 | 0.0075363 | -0.0075528 | 0.0082899 | 0.0067826 | 0.0090435 | 0.006029 | 0.1977 | 0.49634 | 0.028014 | 5268.6451 | 0.0030605 | 0.00069247
Display Cleanliness |  | Pass | L'min (cd/m2) | 1.42 | 180 | 130.31 | 0.19806 | 0.49543 | [5 empty] | 130.41 | 512.3599 | 513.3478 | 131.3485 | 2.4078 | 0.016323 | 0.0074529 | 0.0073306 | 0.01668 | 0.0080637 | 0.0065975 | 0.0087967 | 0.0058645 | 0.19806 | 0.49543 | 0.027104 | 5281.67 | 0.0039283 | 0.00098273
Test Pattern High Contrast Features |  | Pass | L'max (cd/m2) | 452.36 | 195 | 168 | 0.19762 | 0.49546 | [5 empty] | 168.1 | 547.6099 | 548.8843 | 169.6467 | 2.35 | -0.0080619 | 0.007106 | 0.007161 | -0.007688 | 0.0078771 | 0.0064449 | 0.0085932 | 0.0057288 | 0.19762 | 0.49546 | 0.027135 | 5300.9261 | 0.003598 | 0.00044111
Test Pattern Low Contrast Features |  | Pass | AR | 0.075758 | 210 | 216.66 | 0.19708 | 0.49602 | [5 empty] | 216.76 | 583.624 | 584.4207 | 218.005 | 2.4009 | 0.013441 | 0.0071158 | 0.0070208 | 0.013539 | 0.0077228 | 0.0063187 | 0.0084249 | 0.0056166 | 0.19708 | 0.49602 | 0.027702 | 5307.5278 | 0.0028193 | 0.0007795
Test Pattern Grayscale Ramp |  | Pass | LR | 318.5634 | 225 | 278.19 | 0.19657 | 0.49614 | [5 empty] | 278.29 | 619.6261 | 619.9572 | 278.9765 | 2.4001 | 0.013102 | 0.0069951 | 0.0069047 | 0.013097 | 0.0075951 | 0.0062142 | 0.0082856 | 0.0055237 | 0.19657 | 0.49614 | 0.02784 | 5326.8721 | 0.0024127 | 0.00052086
|  |  | Mean JND/GL | 2.3691 | 240 | 355.31 | 0.19622 | 0.49679 | [5 empty] | 355.41 | 655.3697 | 655.4937 | 355.766 | 2.3829 | 0.005831 | 0.0068492 | 0.0068086 | 0.0059565 | 0.0074895 | 0.0061278 | 0.0081704 | 0.0054469 | 0.19622 | 0.49679 | 0.028506 | 5321.4303 | 0.0016804 | 0.00073767
Uniformity (Fill in luminance reading in appropriate cell) |  |  | JND/GL Max Err | 4.407% | 255 | 452.26 | 0.19521 | 0.49814 | [5 empty] | 452.36 | 691.0301 | 691.0301 | 452.3984 | 2.3774 | 0.0034876 | 0.0067549 | 0.0067294 | 0.0037766 | 0.0074024 | 0.0060565 | 0.0080753 | 0.0053836 | 0.19521 | 0.49814 | 0.029921 | 5322.7812 | 0 | 0.0016804
| 22.023 |  | dLL/JND Max Err | 3.798%
22.1 | 21.694 | 19.38
| 19.729 |  | White Point
Unif. Meas. Level | 120 |  | Τ1 (Τ1max) | 0.0039 (0.0077)
MLD | 13.11% |  | Τ2 (Τ2max) | 0.0014 (0.0031)
LUDM | 10.62% |  | Max ΔD65 | 0.026681
Result | Pass |  | Mean CCT | 5299.3336
[10 empty] | Prev Lamb | 0.1 |  |  | GSDF Coefficients | [6 empty] | 110% | 90% | 120% | 80% | [6 empty] | Τ1 | Τ2
[14 empty] | a | -1.3012 | A | 71.4981 |  | JND/GL | 2.3691 | 2.606 | 2.1322 | 2.8429 | 1.8953 | [6 empty] | 0.0038512 | 0.0014093
[14 empty] | b | -0.02584 | B | 94.5931
[14 empty] | c | 0.080243 | C | 41.9121 |  | dGL | 15
[14 empty] | d | -0.1032 | D | 9.8247
[14 empty] | e | 0.13647 | E | 0.28175 |  |  | u' | v'
[14 empty] | f | 0.028746 | F | -1.1878 |  | D65 | 0.19783 | 0.46833
[14 empty] | g | -0.025468 | G | -0.18014
[14 empty] | h | -0.0031979 | H | 0.14711
[14 empty] | k | 0.00012993 | I | -0.017047
[14 empty] | m | 0.0013635